# SP	13549

REVISION:	A

DATE REV:	08/05/04

PROJECT:	MAESTRO

## PROCEDURE

### FOR MAESTRO LENS CAN OPTICAL ALIGNMENT

SP13549

Prepared By:	Michelle Reed

Date:	08/05/04

APPROVALS:

Optical Engineer:	Michelle Reed	Date: 08/05/04

Mechanical Engineer:	Scott Mathews	Date: 08/05/04

Program Scientist:	Jill Bechtold	Date: 08/05/04

Released By:	Greg Landis	Date: 08/05/04

TABLE OF CONTENTS:

Table of Contents								2

Equipment needed								3

Chemicals needed								3

Reference Documents								3

Introduction									3

Temperature Considerations							5

Alignment Location								5

1.0 Alignment of Practice Optics						6

  - Alignment of Test Lenses 1 and 2				6

  - Alignment of CaF2 Lens						8

  - Assembly of Test Lens Triplet					9

  - Oil Check of Assembled Triplet					10

2.0 Mounting of Lenses into Cells						11

2.1 Mounting of Spare Lens 3						11

2.2 Lens 2 through 6 Mounting Procedure				14

- Stack Process								17

  - Triplet								17

  - Lenses 2 and 3							20

  - Registration of Doublet and Triplet				21

  - Addition of Oil							22

- Lens 1 Alignment								22

Appendix A   Tolerances							25

Appendix B   Sylgard Mixing Procedure					26

MAESTRO LENS CAN OPTICAL ALIGNMENT PROCEDURE

Document number SP13549

INTRODUCTION

This procedure outlines the necessary steps for the alignment of the full Maestro Optical Assembly.  Included in this procedure is the mounting of the test optics in the test mounts, mounting of the lenses into their cells, alignment of the lens can, and the alignment of the rest of the instrument.

EQUIPMENT NEEDED

1.Dial indicator accurate to .0001 inch.

2. Dial indicator stands/holders.

3.  Rotation table

4.  Low power laser.

5.  Alignment Telescope

6.  Portable interferometer for final collimator check.

7.  Large flat mirror to place under the test set up to observe wobble.

8.   Suction lifting assembly to lower lenses into mounts (possibly one available to borrow from mirror lab)

- Jig to hold laser and alignment telescope above rotational table to measure beam wander.

- Gauge blocks.

- Paint brushes

- Glue Dispenser

- 10cc and 30 cc barrels for dispenser.

- Low viscosity plunger for barrels.

- Green and purple barrel tips.

- Face masks, respirators.

- Gloves, latex and nitrile.

- Beakers

19. CMM machine available in the Instrumentation shop

CHEMICALS NEEDED

- Acetone for cleaning the cells prior to priming, use gloves while using.

- DOW 92-023 primer or other sylgard acceptable primer, use nitrile gloves and respirator while applying.

- Sylgard 184 epoxy, use gloves while using.

- Dow Corning 760 sealant.

- Dap fun-tak putty or similar to dam up the cell edges and setscrew holes.

- Compressed nitrogen for operation of dispenser.

REFERENCE DOCUMENTS

- DRAWINGS

  - 13473 Lens Can- Lens 2 to 6 Assembly

  - 12871 Lens – Element 2

  - 12872 Lens- Element 3

  - 13198 Lens Can- Lens 2 to 6- Lens 2 Cell

  - 13199 Lens Can- Lens 2 to 6- Lens 2 Cell 2 to 3 Spacer Spool

  - 13200 Lens Can- Lens 2 to 6- Lens 3 Cell

  - 13204 Lens Can- Lens 2 to 6- Lens Cell Spacer Blank

  - 13485 Lens Can- Lens 2 to 6- Lens Clip

  - 13550 Lens Can- Lens 4 to 6 Assembly

  - 12873 Lens- Element 4

  - 12874 Lens- Element 5

  - 12875 Lens- Element 6

  - 13201 Lens Can- Lens 2 to 6- Lens 4 Cell

  - 13202 Lens Can- Lens 2 to 6- Lens 5 Cell

  - 13203 lens Can- Lens 2 to 6- Lens 6 Cell

  - 13194 Lens Can- Lens 1 to 6 Assembly

  - 13472 Lens Can- Lens 1 Stage and Extension Assembly

  - 13474 Lens Can- Lens 1 Stage Centering Adjustment Block

  - 13478 Lens Can- Lens 1 Stage Tip/Tilt Adjuster

  - 13557 Lens Can- Lens 1 Stage Assembly

  - 13196 Lens Can- Lens 1 Stage Shell Front Flange

  - 13205 Lens Can- Lens 1 Focus Stage Shell Stringer

  - 13206 Lens Can- Lens 1 Focus Stage Cross Roller Guide Stringer

  - 13558 Lens Can- Lens 1 Stage Shell Assembly

  - 13195 Lens Can- Lens 1 Stage Shell

  - 13211 Lens Can- Lens 1 Stage Lens 1 Cell

  - 13476 Lens Can- Lens 1 Stage Cross Roller Guide Index Bar

  - 13477 Lens Can- Lens 1 Focus Stage Lens 1 Clip

  - 13197 Lens Can- Lens 1 Stage Lens 1 to 2 Spool Weldment

  - 13560 Lens Can- Lens 1 Stage Lens 1 to 2 Spool Assembly

  - 13220 Lens Can- Lens 1 Stage Lens 1 to 2 Spool Stringer

  - 13486 Lens Can Lens 1 to 2 Spool Spacer

  - 13901 Lens Can- Lens 1 Focus Stage Rear Limit Block

  - 13900 Lens Can- Lens 1 Focus Stage Front Limit Block

  - 13481 Lens Can-Lens 1 Stage Cross Roller Guide Index Bar Nutplate

  - 13902 Lens Can- Lens 1 Focus Stage Cross Roller Guide Preload Plate

  - 13559 Lens Can- Lens 1 stage Lens 1 cell Assembly

  - 13475 Lens Can-Lens 1 stage Shell Back Flange

  - 13592 Lens Can- Lens 1 Stage Shell

Drawing list to be updated on final Lens 1 design completion

- Procedures:

  - SP13879- Installation and Alignment of Crossed Roller Slide Rails Onto L1 Cell.

  - SP14099- System Alignment Procedure

  - SP14109-  Space Frame Alignment Proedure

TEMPERATURE

1.  The optimal alignment temperature for our system is 60 degrees F to ensure acceptable alignment over the full operational range.    If it is not possible to obtain this temperature in the laboratory, the system will have some degradation of resolution at –12 degrees C and lower.

2.  Current work being accomplished in order to be able to cool room 353C to 60 degrees.  The Instrumentation shop CMM machine is in a controllable environment.

Location for alignment of the lenses:

- The alignment of the lenses in their cells will be accomplished in the instrumentation shop using the rotation table for primary alignment with the CMM machine for final alignment and verification.  If extended periods of time are not available on the CMM, Room 353C will be used with the rotation table as the full alignment mechanism.  Preliminary experimentation with the 4” optic, and crude adjustment mechanisms indicated that alignment to within + 0.0005” both on the flatness of the top surface of the mount and the run-out around the edge will be possible.

The rotation table in the laboratory was characterized.  The TIR of the side of the table, 2/3 of the way down from the top, is 0.0021”, however the ending points consistently were less than the starting points by 0.0003”. The top of the table wedged with a TIR 0.00185”, but it was repeatable by position to 0.0001”.  Backlash on this table was significant, but could be controlled with the side tightening screws.  Also when reversing directions you came back to the same point when reversing back to the original direction.  With some care, I believe we can use this table. Shimming can be performed to ensure the levelness of our test set up.

  - ALIGNMENT OF THE PRACTICE OPTICS:

Gloves should be worn during all optical alignment procedures, to prevent contamination.

Prior to any alignment with the MAESTRO optics, we will use the 3 small test optics to verify the oiling and cell mount procedure.  These elements are approximately 50mm diameter and we can use the procedure as listed for lenses 4,5, and 6.  We will also obtain a full size lens cell for use with the spare lens 3 to verify the lens mounting procedure on a full size sample.

1.1 Test Lenses 1 and 2

- While wearing gloves to keep the shim stock clean, cut the appropriate thickness of plastic shim stock to the correct inner diameter shelf size.  Make sure you subtract 0.001 inches thickness to account for the tape thickness that will be used to adhere the shim to the lens cell.  Keep the shim clean, as using acetone will strip some of the color applied, which will also affect its thickness.

- Place the lens cell for one of the outer lenses for the triplet on the table so that it is roughly centered using a micrometer and lock the cell in place, as shown in figure 1.     Place the dial indicator on the top of the lens cell, and adjust the cell until it is level to within tolerance as per Table 1, appendix A.  Add shims as necessary.

Figure 1

- Move the dial indicator to the side of the lens cell, as shown in figure 2 and measure the runout of the lens cell.  Adjust the lens cell from side to side, until the total runout of the cell is within tolerance per Table 1.  Lock the lens back into place.

Figure 2

- Check the wedge on the top surface again to make sure no wedge has been introduced.  Adjust the wedge and centration again if necessary.  Record the final errors in the alignment in Table 1.

- Paint the inside of the lens cell surface with primer and allow to dry for a minimum of 90 minutes.  Insert the lens pad shim, and secure with double sided tape.

  - Thickness of double sided tape: 0.001”

  - Thickness of shimstock used for pad:  0.002”

- Place the lens in the cell with the plane parallel side down, and roughly center it using a micrometer.  Since no set screws would fit in the test optic cell, and the force applied by the dial indicator is stronger than the frictional force of the lens to the shelf, exact centering of the lens will not be accomplished.  Cut shrink wrap into small strips, layer them together to provide resistance to flex, and place at 4 points around the outside of the lens at 90 degree increments.  This will provide equal resistance around the outside of the lens cell, which in turn will roughly center the lens.

NOTE:  The second test cell was centered using very fine shims made from wood, to provide the centering, as the sylgard did not cure where in contact with the shrink wrap.

- Place the dial indicator at the top edge of the lens and measure the wedge of the lens cell.  Adjust adding shim stock as necessary between the lens and the cell to level the lens to tolerance per Table 1.

- Check the runout on the centration.  Adjust wedge and centration as necessary.  Record the final errors.

NOTE: Were not able to obtain the required runouts on the centration or wedge due to irregularities at the point of measurements, difficulty in room to add shims, and drag noticed with the dial indicator.

- Mix the appropriate amount of Sylgard, evacuate the majority of bubbles.  Do not pump the sample for longer than 45 minutes.

- Hook up the automatic dispenser to the compressed nitrogen bottle to provide the appropriate pressure.  Using the dispenser, apply the Sylgard 184 layer carefully around the full edge of the lens with a continuous bead, slowly adjusting the pressure on the dispenser if the sylgard is dispensing too slowly or too swiftly.

- Allow the Sylgard to set.  It will take 4 days for the Sylgard to set sufficiently for moving the lens off of the table.  Once the SYLGARD layer has set for at least 7 days, seal it with an oil resistant seal such as RTV 730 to seal all bond lines that may come in contact with the oil.

- Mount test lens 2 using the same method as listed above.

NOTE:  Instructions for the primer state that you should not let the primer set more than 6 hours prior to gluing, so primering of the cell should not occur before you are ready to align the lens in the cell and proceed immediately to the glue stage.

NOTE:  Encounters upon gluing.  Glue leaked out under lens.  Gap under lens less than 8 thou.  Sylgard did not cure where it came in contact with the shrink wrap.  After 5 days it was still very glue-like.

  - MOUNTING OF THE TEST CALCIUM FLOURIDE LENS:

A different process is used for this lens because there is not a ledge on the cell to rest the lens against.

- Measure the height of the outer adapter ring with the lens cell, at the top edge of the interior edge of the lens cell.  Next measure the height of the lens on the ball bearings, with the ring systems.  Compare the two measurements and add shims as necessary to make the two measurements equal.

- Place the standoff test ring, with internal ring, o-ring seal, and ball bearings on the rotation table, and roughly center it using a micrometer.  Lock it in place.

- Place the dial indicator at the outer edge of the raised alignment jig, and level the assembly to within tolerance per Table 1, shimming as necessary.  See Figure 3.

Figure 3

- Next place the dial indicator at the outer edge of the alignment jig and center it to within tolerance per Table 1.  Recheck the level.  Repeat centering and leveling until within specification.  Clamp the test set up in place.

- Place the CaF2 lens cell on top of the jig and place the dial indicator on the top surface.  Measure the total indicator run out.  Wedge as necessary to level the cell.  Measure the height of the top lens mounting edge.

- Place the lens onto the set up, and roughly center it using a micrometer.  Place the dial indicator on to the top of the lens, outside the clear aperture.  Measure the total run out.  Level the lens by shimming until it is within tolerance per Table 1, wedging under the o-ring as necessary.    Measure the height of the top edge of the lens, once the lens is roughly centered and flat.  This height will be verified at several of the positions around the outside edge of the lens, to check the wedge of the lens.

- Compare this height to the cell height.  If the cell height is too low, shim it higher, until the heights are the same.  If the lens height is too low, remove the lens and raise the inner ring until it is at the appropriate height for the lens.  Apply primer to the cell wall at this point.  Return the lens to the assembly and repeat steps 6 and 7 until the height of the lens and cell are correct for gluing, and have been properly leveled.

- Again, there is no mechanism to appropriately center the test lens, shim until it is roughly centered. Record the final wedge and decentration of the test CaF2 lens and cell:

- Apply the sylgard using the dispenser as outlined in step 10 of the previous section, until the appropriate thickness of the glue layer is obtained.  Verify that no wicking of the glue occurs past the o-ring seal.   Allow the glue to cure in place.  Once the glue has cured well (several days potentially), remove the test adapter plates and the o-ring seal.  Allow sylgard to cure for a minimum of 7 days.

#### 1.3 ASSEMBLY OF THE TEST LENS TRIPLET

Once the individual lens have cured sufficiently, the test triplet assembly will be assembled as follows:

- Place the CaF2 lens assembly on the rotational table with the lens on the top side, and roughly center it using a micrometer.  Secure the cell in place so that it will not move unintentionally, but can still be shifted for centration.

- Place a dial indicator on the outer top edge of the cell and measure the runout.  Using shim stock, level the cell until it is within the tolerance.

- Place the dial indicator on the outer top edge of the cell, and center the cell until it is within tolerance.  Lock it in place.

- Verify the levelness of the cell again.  Repeat steps 2 and 3 until proper level and centration are obtained.

- Place the o-rings in place on the lens 1 cell.

- Place the lens 1 cell on top of the centered cell with the lens facing down, and roughly center it.

- Place the dial indicator on the top edge of the cell and measure the run out.  Do not shim the cell, as it could compromise the function of the o-ring seal.  Record the wedge.

- Place the dial indicator on the outer edge of the cell and measure the total run out.  Adjust the centration of the lens cell until the total runout is within tolerance.  Lock into place.

- Lock the two lenses in place together.

- Flip the assembly over so that lens 1 is on the bottom, and lens 2 is on the top.  Roughly recenter the lens assembly on the table using the micrometer.  Lock in place.

- Place the dial indicator on the top of the outer edge of the assembly and measure the wedge.  Add shims under the whole assembly to eliminate this wedge.  Record final errors in decentration and wedge.

- Place the next o-ring seal on the lens 2 cell.

- Place the test lens cell 2 onto the assembly with the lens facing down.

- Roughly center the lens 2 cell using a micrometer.  Place the dial indicator on the outer top surface of the cell and measure the total run-out on the cell. Do not try to level it at this point. As it may interfere with the operation of the o-rings.  Record the runout.

- Place the dial indicator on the outside edge of the top cell and measure the total runout.  Adjust the lens 2 cell until the appropriate tolerance is obtained.  Lock the cell into position with the other two lenses.  Record its decententration.

#### 1.4 OIL CHECK OF ALIGNED SYSTEM

- Verify that the triplet is sealed.

- Install bleed off bladder. (Brake bleed?)

- Inject oil into the system, using the dispenser gun.

- Bleed off any air.

- verify system is filled with oil.

- Monitor the system for leaks.

You have now completed the alignment of the test optics and are ready to move on to MAESTRO system optics.

2.0 MOUNTING OF LENSES IN CELLS

- Place a reference flat on the rotation table.  This reference flat should have a diameter slightly less than the inner diameter of the lens cells, to allow room for the gauge blocks.  The exact diameter will vary with what is available from the mirror lab.  Roughly center the reference flat using a micrometer.

- Place the dial indicator on the very outer top edge of the reference flat, and measure the total runout.  Shim the mirror as necessary to ensure a level surface.

- Mount a laser above the table, using a jig that will allow for the mounting of the laser and a turning flat to allow for observing the reflection of the laser off the lower mirror by an alignment telescope mounted sideways (See figure 1) or observing the reflection on the wall/ceiling to observe movement with rotation.

- This will complete the set up of the alignment jigs.

#### 2.1 MOUNTING OF THE SPARE LENS 3

Verify the full size lens alignment with the spare lens 3 assembly as follows:

- Place gauge blocks on the rotational table so they are equidistant from the center of rotation of the table.

- Place the lens cell for the spare lens 3 on gauge blocks, facing up. Roughly center the cell using a caliper.

- Place the dial indicator on the outer top surface of the cell and measure the total run-out of the top surface of the cell (Measurements every 20 degrees initially provides a good characterization).   Note the high and low spots and shim as necessary until the cell is level to well within tolerance of Table 1. Adjust an outer edge holding jig so that you can use set screws to move the cell radially without interfering with the reference edge of the cell.  Adjust until it is barely holding the cell, but not so that it will deform the surface.  Check the levelness of the cell after holding in place.

- Move the dial indicator so that it resting against the reference surface on the outside edge of the cell.  Check the centration.  Using the set screws on the holding jig, adjust the lens until it is centered on the rotation of the table within tolerance per Table 1.  Lock the cell into place, so no accidental movement of the cell is possible.

- Verify the centration and levelness of the lens cell.  Adjust as necessary.

- When you are ready to begin aligning the lens 3 into the cell, paint the inside of the lens cell surface with the 92-023 primer.  Safety precautions should be taken. Once primer is applied, the lens must be aligned.  The manufacturer recommends that the adhesive be applied over the primer within 6 hours of primer application.  Tests performed by the Steward Observatory AO group indicate that overnight drying of the primer does not affect the performance of the adhesive strength.

- Place the appropriate precut lens pad into the cell and tape into place.

- Place the spare lens 3 into the locked down cell using the suction holder.  Place the lens so that the concave surface is facing up. See Figure 4.

Figure 4

- Roughly center the lens into the holder using shims.  Once roughly centered, remove the shims.

- Place the dial indicator on the top of the outer edge of the lens, and measure the levelness of the lens.  Shim under the lens until the lens is level to within the tolerance of Table 1.  Move the set screws in until they are barely resting against the lens edge but are applying no stress to the lens.

- Move the dial indicator to the outside edge of the lens, and measure the centration.  Using the set screws, slowly adjust the set screws to center the lens, but do not stress the lens or the cell.  Once the acceptable centration is achieved, adjust the set screws until they are stably holding the lens in place.

- Set another dial indicator on the top ledge of the lens and check the levelness of the lens.  Adjust the leveling shims as necessary.

- Direct the laser through the lens and monitor the positioning of the reflection using an alignment scope, or reflection on a wall/ceiling.  Monitor movement of the laser position on the mirror as the table is rotated 360 degrees.  Adjust the centration and the wedge until there is no observable movement of the laser positioning.  Check the alignment with the dial indicators.

- Tighten the setscrews and verify the lens positioning again.  Repeat steps as necessary.  Remove temporary shims.

- Once the lens is securely locked in place, carefully apply putty to the edge of the cell shelf so that it does not come in contact with the clear aperture of the lens, but provides for a barrier to prevent leakage of the sylgard.

- Transfer the lens to the CMM table.  Characterize the lens and the cell to make sure that they are concentric and level.  Adjust positioning of the lens as necessary to increase accuracy.

- Mix a SMALL amount of sylgard and apply in a continuous bead around the entire cell width.  Apply enough sylgard so that it can wick onto the shelf upon which the lens sits, but so that there is not enough to leak out the other side.  Allow to set sufficiently to aid in sealing the possible leak points. Cure and label a witness sample.

- After the small amount of sylgard applied has set partially, mix remaining appropriate amount of Sylgard per directions, and evacuate bubbles.  Load dispenser with sylgard, and using the automatic dispenser apply the appropriate amount of Sylgard 184 in a continuous bead around the entire cell.  Continue evenly around cell to the desired sylgard width, and let set completely.  Cure and label a witness sample.

- Remove putty and clean surfaces affected.  Loosen set screws slightly so that no stress will be introduced to the lens.

- Mark the outside edge of the cell with an arrow to indicate the initial light path direction, and any remaining high or low points for the lens/cell combination, both in parallelism, and in decentration of the lens.

- Modify this procedure as necessary based on the obtained information of this test mounting.

- Continue mounting of the balance of lenses as per the following section.

2.2 LENS 2-6 MOUNTING PROCEDURE

NOTE:  Element tilts for lenses 2 and 3 are critical with the tolerancing.  Make sure these alignments are performed extremely carefully.

- Place gauge blocks on the rotational table so they are equi-distant from the center of the table, and so that the lens 6 cell may be placed on the gauge blocks without interfering with the cut out.

- Place the lens 6 cell on gauge blocks, so that the lens may be placed in on the ledge. Roughly center the cell using a caliper.

- Place the dial indicator on the outer top surface of the cell and measure the total run-out on the cell.   Measure the high and low spots and shim as necessary until the cell is level to well within tolerance per Table 1.

- Adjust an outer edge holding jig so that you can use set screws to move the cell radially without interfering with the reference edge of the cell.  Adjust until it is barely holding the cell.  Check the levelness of the cell after holding in place.

- Move the dial indicator so that it resting against the reference surface on the outside edge of the cell.  Check the centration.  Using the set screws on the holding jig, adjust the lens until it is centered on the rotation of the table within tolerance of Table 1.  Lock the cell into place.

- Verify the centration and levelness of the lens cell.  Adjust as necessary.

- When you are ready to begin aligning lens 6 into the cell, paint the inside of the lens cell surface with the 92-023 primer.

- Place the appropriate precut lens pad into the cell and tape into place, using two sided tape.

- Place lens 6 into the locked down cell using the suction holder.  Place the lens so that the concave surface is facing up.  Roughly center the lens into the holder using shims.  Once roughly centered, remove the shims.  See Figure 5.

Figure 5

- Place the dial indicator on the top of the outer edge of the lens, and measure the levelness of the lens.  Shim under the lens until the lens is level to within tolerance Table 1.  Move the set screws in until they are barely resting against the lens edge but are applying no stress to the lens.

- Move the dial indicator to the outside edge of the lens, and measure the centration.  Using the set screws, slowly adjust the set screws to center the lens, but do not stress the lens or the cell, until the acceptable centration is achieved to within tolerance of Table 1.  Adjust the set screw until they are stably holding the lens in place, but applying no stress to the lens.

- Set another dial indicator on the top ledge of the lens and check the levelness of the lens.  Adjust the leveling shims as necessary.

- Direct the laser through the lens and monitor the positioning of the reflection using an alignment scope.  Monitor movement of the laser position on mirror as the table is rotated 360 degrees.  Adjust the centration and the wedge until there is no observable movement of the laser positioning.

- Tighten the setscrews and verify the lens positioning again.  Repeat steps as necessary.  Once lens is locked into place, carefully remove temporary shims, and re-verify locked down position.  Carefully apply putty to the edge of the cell shelf so that it does not come in contact with the clear aperture of the lens, but provides for a barrier to prevent leakage of the sylgard.

- Gently transfer the lens and cell to the CMM table, and reverify positioning of the lens with respect to the cell.  Adjust positioning of the lens with the cell if necessary, to increase the positioning accuracy to within + 0.0004”.

- Mix a small amount of sylgard and apply in a continuous bead around the entire cell width.  Apply enough sylgard so that it can wick onto the shelf upon which the lens sits, but so that there is not enough to leak out the other side.  Allow sylgard to set sufficiently to aid in sealing the possible leak points. Cure and label witness sample.

- Mix Sylgard per directions, and evacuate bubbles.  Load dispenser with sylgard, and using the automatic dispenser apply the appropriate amount of Sylgard 184 in a continuous bead around the entire cell.  Continue evenly around cell to the desired sylgard width, and let set completely. (Minimum of 7 days).  Cure and label witness sample.

- Remove putty and clean surfaces affected.  Loosen set screws slightly so that no stress will be introduced to the lens.

- Mark the outside edge of the cell with an arrow to indicate the initial light path direction.  In addition mark low and high points and amounts of the decentration and wedge so that the best configuration can be obtained during the stack process. These marks should be on the edge of the cell for easy reference.

- Mount lens 4 using this same procedure.  Lens 4 should be placed into the cell with the concave surface facing up.  See Figure 6.

Figure 6

- Mount lens 2 using this same procedure.  Lens 2 should be placed into the cell with the plano side up.  See Figure 7.

Figure 7

- Mount lens 3 using this same procedure.  Lens 3 should be placed into the cell with the concave surface up.  See Figure 8.

Figure 8

- Mount Lens 5 using this same procedure.  Lens 5 should be placed into the cell with the smaller radius resting against the ledge.  See Figure 9.

Figure 9

- The mounting of the lenses into their cells is now complete.  The next step is to start the stack process.

## 3.0 START OF THE STACK PROCESS

- Schedule a block of time with the CMM machine at the Instrument lab.  A minimum of 2 months is desired.  Transfer the rotation table to this location to provide for double checks.

- Adjust the thermal control of the laboratory to 60 degrees, if possible.

  - Triplet

- Place Lens 6 on the CMM table on gauge blocks, so that lens 5 can be stacked on top of it.  Fix lens 6 into its position on the CMM table.  This position will provide the main reference for stacking of the triplet. Do not move its position or re-home the CMM machine until the alignment of the triplet is complete.  See figure 10.

Figure 10

- Quantify the positioning of the mount of lens 6, and the edge of the optic mounted inside.  Adjust as necessary to un-wedge any left over residual.  Fully characterize the positioning of both the lens and the cell., verifying that the plane of the edge of lens 6 is parallel to the plane of the CMM table, and that the optical axis defined by lens 6 is perpendicular to the CMM table.

- Once this lens has been fully characterized in its fixed position, place the appropriate spacer with lens 5 on top of it.  The smaller radius of lens 5 will be facing towards lens 6 (This is the shelf side of the mount.).  Do not yet install the o-rings as they could lead to errors in spacing.  See figure 11.

Figure 11

- Center lens 5 horizontally so that it is concentric with lens 6 to within the alignment tolerances as specified in table 1.

- Measure the position of the top edge surface of lens 5 and compare it to the calculated required position.  Check the measurements for wedge and spacing.

- Modify the spacer by approximately ¾ of the required correction to obtain the proper spacing.

- Remount lens 5 with the modified spacer.  Remeasure and repeat steps 7-8 until lens 5 is at the proper location, to within the acceptable tolerance.

- Once appropriate spacing is obtained, index mark orientation of the spacer/lens assemblies.  Add o-rings to their positions, and clamp the pieces together to the appropriate torque

- Verify spacing of the system.

- Repeat steps 6-9 until alignment is regained.

- Direct a laser through the elements and monitor the positioning of the reflection using an alignment telescope. Monitor movement of the laser position on mirror as the table is rotated 360 degrees.  Adjust the centration and wedge until there is no observable movement of the laser positioning.  Regain appropriate torque.

- Add lens 4 and the 4-5 spacer onto lens 5 with the curved surface facing lens 5.  See figure 12.

Figure 12

- Repeat steps 6-9, comparing lens 4 position with lens 6 position.

- Once appropriate spacing is obtained, index mark orientation of the spacer/lens assemblies.  Add o-rings to their positions, and clamp the pieces together to the appropriate torque.

- Verify spacing of the system.

- Repeat steps 6-9, comparing lens 4 position with lens 6 position

- Once proper alignment has been obtained, bolt lens 4 together with lens 5-6.

- Direct a laser through the elements and monitor the positioning of the reflection using an alignment telescope. Monitor movement of the laser position on mirror as the table is rotated 360 degrees.  Adjust the centration and wedge until there is no observable movement of the laser positioning.  Regain appropriate torque.

- Once satisfied with the alignment, make sure that the elements are well bolted so that no movement can occur with respect to each other.

NOTE:  If the cut away of lens 6 interferes with the stacking on the table, we will reverse the order of the previous alignment.

NOTE:  Interferometric testing is not possible at this point without the design and fabrication of a holographic null lens.

3.2 LENSES 2 AND 3

NOTE:  Element tilts for lenses 2 and 3 are critical for tolerancing.  Make sure these alignments are performed extremely carefully.

After Lenses 2 and 3 are mounted individually into their cells using the procedure listed in the previous portion of this document, they will be aligned to each other as follows:

- Place Lens 2 on gauge blocks on the CMM table so that the curved surface is facing down.  Characterize the position of the lens and cell.  See Figure 13.

Figure 13

- Define the optical axis from the characterization measurements, and shim until the optical axis is perpendicular to the table, and the plane of the lens is parallel to the table.

- Lock the lens into position so that it can become the reference position.  Recharacterize the position.  Do not move its position or re-home the CMM machine until the alignment of lens 2 to lens 3 is complete.

- Place lens 3 and the spacer between lens 2 and 3 on top of lens 2 with the concave surface facing down.  See Figure 14.  Adjust lens 3 horizontally until the two lenses are concentric within the tolerances listed in Table 1.  Measure the height and levelness of lens 3 and compare it to the required positioning.

figure 14

- Modify the spacer by approximately ¾ of the required correction to obtain the proper spacing.

- Repeat steps 3 and 4 until the proper positioning of lens 3 is obtained.

- Bolt the two lenses together and verify alignment.

- Transfer the lenses to the rotation table and center them to within tolerance.  Direct a laser through the lenses and monitor the laser positioning through an alignment telescope while the lens is rotated 360 degrees.  Adjust positioning if necessary.

3.3 REGISTRATION OF DOUBLET AND TRIPLET

- Return the lens 2 and 3 to stack position on the CMM table with lens 2 down.

- Lens 3 will be held in position at this point, so that the triplet will be movable if necessary.  The position of lens 3 will be quantified at this point to be used as the reference point.

- Place the spacer for lens 3 to 4 and the triplet with lens 4 down on top of the doublet assembly and guide the triplet into position with push/pull screws, until it is concentric with the positioning of the lens 2/3 assembly, within the toleralnces of table 1.

- At this point, the top of lens 4 will not be accessible, and spacing measurements will have to occur between lens 3 and lens 6.  Measure the spacing between the referenced lens 3 and lens 6.  Compare the positioning to the required spacing of these two elements.

- Adjust the spacer thickness for the lens 3 to 4 spacer by ½ the required correction, and repeat the process to ensure that the appropriate spacing increment has been achieved.  Continue measurements and adjustment of the spacer until the desired spacing is obtained.

- Check the positioning of the lens assembly for perpendicularity of the lens optical axis with the referenced lens 3 position, and the concentricity of the elements.

- Lock the two subassemblies together and transfer it to the rotation table.  Center the system on the table, to within the tolerance of table 1.

- Observe the laser through the system with the alignment scope.  Verify the alignment and adjust as necessary.

- Perform a star test or interferometrically test lenses 2-6 to verify alignment.  Final adjustments will be made as necessary.  This test should be accomplishable with a precision flat to auto collimate the beam.  Oil must be added to the system to perform the optical verification.

###### 3.4 ADDITION OF OIL

This section is to be completed after definition of assembly is finalized.

4.0 LENS 1 ALIGNMENT

For the alignment of the cross roller slides to the cell assembly please refer to document number SP13879.  Below begins the alignment procedure for the installation of lens 1 to its cell, and the alignment of the completed assembly to the rest of the lens can.

- With the lens one cell on its side, as it will be mounted in the space frame, index mark 12 positions equally spaced around the outside of the inner cell for lens 1.  Using the CMM, measure these 12 positions.

- From these 12 positions, define the center position of the cell.

- At eight translation positions along the cross roller slides measure the positioning of the 12 points of the cell with the CMM machine.  Using the CMM define the movement centerline and error of this line.  This will become our optical axis of this assembly.

- Measure the front surface of the lens cell to define the plane.

- Measure the perpendicularity of the lens 1 faceplate plane and the optical axis plane.  Note any deviations.

- Lens 1 will have to be mounted perpendicular to the calculated optical axis, and on center with it to within the tolerances of table 1.  Turn the assembly so that it is vertical.  Shim the assembly until the top surface of the inner lens 1 cell is perpendicular to the optical axis.  Using the CMM, shims as necessary and the side set screws, align lens 1 until it is in position with the optical axis of the lens 1 assembly to within the tolerances specified in table 1.

- It is anticipated that the way we have specified the mechanical mounting of the cross roller slides, that the front surface of lens 1 should be perpendicular to the optical axis, and very little shimming should be necessary between lens 1 and lens 1 cell.

- Once the position of lens 1 is correct in the cell and is firmly being held in place with the setscrews, turn the lens over and apply a very fine layer of putty against the shelf, but outside of the clear aperture of lens 1.  This will prevent wicking of the sylgard onto the surface of lens 1 if shims are too great.

- Turn the lens back over and reverify the positioning of the lens.  Adjust if necessary.

- Mix a small amount of sylgard and apply it in a thin bead around the cell.  This will be used to wick through the shelf and provide an additional barrier to the full sylgard amount.  Cure and label a witness sample

- Once the first amount of sylgard has set so it will no longer flow, mix the remaining amount of sylgard and apply it to fix the lens within the cell.  Cure and label a witness sample.    While the sylgard is setting, it is important not to move the lens.  It will take 4 days for the sylgard to harden to the point that the lens will be movable from its location.

- Return the lens 2-6 assembly to the rotation table with lens 6 down.  Center the assembly to within the tolerance of table 1.  Next, attach the spool assembly that connects the lens 1 assembly to lenses 2-6.  Center this with the lens 2-6 assembly to within the mechanical tolerances of the spool assembly.

- Attach the lens 1 assembly and center the inner cell with respect to the lens 2-6 assembly, within the tolerances listed in table 1, using the push-pull tip-tilt screws on the lens 1 assembly, while monitoring a dial indicator that is recessed down to the edge of the lens 1 inner cell. Observe a laser with the alignment scope, and make adjustments until there is no wobble of the laser viewed through the lens system, and the dial indicator indicates they are concentric and level.  Check the position of lens one on the cross roller slides to ensure that the position is centered on the available movement of the lens.

- Transfer the entire set up back onto the CMM machine, and measure the distance between lens 1 and lens 6.  Compare this to the required separation.  The size of the spool should be such that no adjustments should be necessary, however, if proper alignment cannot be accomplished while maintaining lens 1 near the center of movement on the rail, adjust the spool size until the desired separation is achieved, and lens 1 is still near it’s center of movement.

- Fix the full collimator together.

- Transfer the lenses to the rotation table and center them to within tolerance.  Direct a laser through the lenses and monitor the laser positioning through an alignment telescope while the lens is rotated 360 degrees.  Adjust positioning if necessary.

- Move the entire assembly to an interferometer, and check the image quality.  Adjust lens 1 if necessary to improve image quality.  If no improvement is obtainable, and the image quality is poor, earlier steps of this procedure will have to be repeated, until collimator is fully aligned to optimum performance.  Reference to the tolerancing done in Zemax should provide an indication on possible starting locations for locating the source of the misalignment.

- Continue on the alignment of the system using procedures:

SP14098- Prism alignment procedure

SP14099- System alignment procedure

SP14109- Space frame construction

Sp???- - Grating alignment procedure

APPENDIX A

TABLE OF FINAL TOLERANCES

| Description | #### Theory Inches | #### Actual Inches (TIR) |
| --- | --- | --- |
| Lens 4 test optic- Marked lens 1 | 2.0 | 1.998 |
| Lens 6 test optic- Marked lens 2 | 2.0 | 1.999 |
| Lens 5 test optic-CaF2 | 2.0 | 1. |
| Error in final centration of the test lens 1 cell | +0.0005 | 0.0007 |
| Error in final wedge of the test lens 1 cell | +0.0005 | 0.0015 |
| Error in final centration of the test lens 1 | +0.0005 | 0.00485 |
| Error in final wedge of the test lens 1 | +0.0005 | 0.003 |
| Error in final centration of the test lens 2 cell | +0.0005 | 0.0010 |
| Error in final wedge of the test lens 2 cell | +0.0005 | 0.0009 |
| Error in final centration of the test lens 2 | +0.0005 | 0.0017 |
| Error in final wedge of the test lens 2 | 0.0005TIR | 0.0005 |
| Error in final centration of the test CaF2 cell: | +0.0005 |  |
| Error in final wedge of the test CaF2 cell | 0.0005TIR |  |
| ##### Error in final centration of the test CaF2 lens | +0.0005 |  |
| Error in final wedge of the test CaF2 lens: | 0.0005TIR |  |
| Error in final centration of the test lens 1 cell in assembly: | +0.0005 |  |
| Error in final wedge of the test lens 1 cell in assembly: | 0.0005TIR |  |
| Error in final wedge of triplet assembly: | +0.0005 |  |
| Error in final centration of the test lens 2 cell in assembly: | +0.0005 |  |
| Error in final centration of lens 2 cell: max to min | +0.0005 |  |
| Error in final wedge of lens 2 cell: max to min | +0.0005 |  |
| Error in final centration of lens 2: | +0.0005 |  |
| Error in final wedge of lens 2: | +0.0005 |  |
| Error in final centration of lens 3 cell: max to min | +0.0005 |  |
| Error in final wedge of lens 3 cell: max to min | 0.0005TIR |  |
| Error in final centration of lens 3: | +0.0005 |  |
| Error in final wedge of lens 3: | 0.0005TIR |  |
| Error in final centration of lens 4 cell: max to min | +0.0005 |  |
| Error in final wedge of lens 4 cell: max to min | 0.0005TIR |  |
| Error in final wedge of lens 4: | 0.0005TIR |  |
| Error in final centration of lens 4: | +0.0005 |  |
| Error in final wedge of lens 6 cell: max to min | 0.0005TIR |  |
| Error in final centration of lens 6 cell: max to min | +0.0005 |  |
| Error in final wedge of lens 6: | 0.0005TIR |  |
| Error in final centration of lens 6: | +0.0005 |  |
| Error in final wedge of Lens 5 cell: | 0.0005TIR |  |
| Error in final centration of Lens 5 cell: | +0.0005 |  |
| Error in final wedge of Lens 5: | 0.0005TIR |  |
| Error in final centration of Lens 5: | +0.0005 |  |
| Actual spacer thickness used between lenses 5 and 6 |  |  |
| Actual spacer thickness used between lenses 4 and 5 |  |  |
| Actual spacer thickness used between lenses 2 and 3 |  |  |
| Actual spacer thickness used between lenses 3 and 4 |  |  |

APPENDIX B

Sylgard mixing procedure