# SP	13549

REVISION:	A

DATE REV:	08/05/04

PROJECT:	MAESTRO

## PROCEDURE

### FOR MAESTRO ALIGNMENT

SP13549

Prepared By:	Michelle Reed

Date:	08/05/04

APPROVALS:

Optical Engineer:	Michelle Reed	Date: 08/05/04

Mechanical Engineer:	Scott Mathews	Date: 08/05/04

Program Scientist:	Jill Bechtold	Date: 08/05/04

Released By:	Greg Landis	Date: 08/05/04

MAESTRO OPTICAL ALIGNMENT PROCEDURE

Document number SP13549

- Scope

This procedure outlines the necessary steps for the alignment of the full Maestro Optical Assembly.  Included in this procedure it the mounting of the test optics in the test mounts, mounting of the lenses into their cells, alignment of the lens can, and the alignment of the rest of the instrument.

EQUIPMENT NEEDED:

1.Dial indicator accurate to .0001 inch.

2. Dial indicator stands/holders.

3.  Air bearing for stacking

4.  Laser.

5.  Alignment Telescope

6.  Portable interferometer for final collimator check.

7.  Large flat mirror to place under the test set up to observe wobble.

8.   Suction lifting assembly to lower lenses into mounts (possibly one available to borrow from mirror lab)

- Jig to hold laser and alignment telescope above rotational table to measure beam wander.

- Gauge blocks.

- Paint brushes

- Glue Dispenser

- 10cc and 30 cc barrels for dispenser.

- Low viscosity plunger for barrels.

- Green and purple barrel tips.

- Face masks

- Gloves, latex and nitrile.

- Beakers

CHEMICALS NEEDED:

- Acetone for cleaning the cells prior to priming.

- DOW 92-023 primer or other sylgard acceptable primer.

- Sylgard 184 epoxy.

- Dow Corning 760 sealant.

- Dap fun-tak putty or similar to dam up the cell edges and setscrew holes.

- Compressed nitrogen for operation of dispenser.

TEMPERATURE

1.  The optimal alignment temperature for our system is 60 degrees F to ensure acceptable alignment over the full operational range.    If it is not possible to obtain this temperature in the laboratory, the system will have some degradation of resolution at –12 degrees C and lower.

2.  Current work being accomplished in order to be able to cool room 353C to 60 degrees.  The large CMM machine is in a controllable environment.

Location for alignment of the lenses:

- The alignment of the lenses in their cells will be accomplished in Room 353C with the rotation table.  Preliminary experimentation with the 4-inch optic, and crude adjustment mechanisms indicated that alignment to within +/- ½ thou both on the flatness of the top surface of the mount and the run-out around the edge was possible.

- If we are unable to obtain our desired accuracies in the lab with the larger optical cells, we will switch to the air-bearing table in the small mirror lab. If the air-bearing table located in the small mirror lab is required for alignment, it must be available to us for an extended period of time.  The maximum diameter of optics that will not have a height requirement is 13inches.  Our outer cell diameter is 12.875, so we should barely clear this requirement.  The only exception will be the set hardware that will hold the cells in place once they are aligned.

The rotation table in the laboratory was characterized.  The TIR of the side of the table, 2/3 of the way down from the top, is 2.1 thou, however the ending points consistently were less than the starting points by .3thou.  The top of the table, wobbled 1.85 thou, but it was repeatable by position to .1 thou.  Backlash on this table was significant, but could be controlled with the side tightening screws.  Also when reversing directions you came back to the same point when reversing back to the original direction.  With some care, I believe we can use this table. Shimming can be performed to ensure the levelness of our test set up.

The air-bearing table has also been characterized for backlash, TIR, and surface flatness.  The surface flatness is less than .3 thou total error around edge of surface.  The TIR is 3.6 thou that is consistently repeatable. +/- .1 thou.  The end point consistently returns to its starting point.  Backlash is repeatable 0.1 thou.

ALIGNMENT OF THE PRACTICE OPTICS:

Prior to any alignment with the MAESTRO optics, we will use the 3 small test optics to verify the oiling and cell mount procedure.  These elements are approximately 50mm diameter and we can use the procedure as listed for lenses 4,5, and 6.  We will also obtain a full size lens cell for use with the spare lens 3 to verify the lens mounting procedure on a full size sample.

Lens 4 test optic- Marked lens 1- Actual diameter = 1.998 inches

Lens 6 test optic- Marked lens 2- Actual diameter = 1.999 inches

Lens 5 test optic-CaF2- Actual Diameter = 1.970 inches

- While wearing gloves to keep the shim stock clean, cut the appropriate thickness of plastic shim stock to the correct inner diameter shelf size.  Make sure you subtract 2 thousandths of an inch thickness to account for the tape thickness that will be used to adhere the shim to the lens cell.  Keep the shim clean, as using acetone will strip some of the color applied.

- Place the lens cell for one of the outer lenses for the triplet on the table so that it is roughly centered using a micrometer and lock the cell in place.     Place the dial indicator on the top of the lens cell, and adjust the cell until it is level to within +/- .5 thou, adding shims as necessary.

- Move the dial indicator to the side of the lens cell, and measure the runout of the lens cell.  Adjust the lens cell from side to side, until the total runout of the cell is within .5 thou.  Lock the lens back into place.

- Check the wedge on the top surface again to make sure no wedge has been introduced.  Adjust the wedge and centration again if necessary.  Record the final errors in the alignment.

Error in final centration of the test lens 1 cell: 0.7 thou max to min TIR

Error in final wedge of the test lens 1 cell: 1.5 thou max to min at diameter 2.5 inches

- Paint the inside of the lens cell surface with primer and allow to dry for a minimum of 90 minutes.  Insert the lens pad shim, and secure with double sided tape.

  - Thickness of double sided tape:  1 thou.

  - Thickness of shimstock used for pad:  2 thou.

- Place the lens in the cell with the plane parallel side down, and roughly center it using a micrometer.  Since no set screws would fit in the test optic cell, and the force applied by the dial indicator is stronger than the frictional force of the lens to the shelf, exact centering of the lens will not be accomplished.  Cut shrink wrap into small strips, layer them together to provide resistance to flex, and place at 4 points around the outside of the lens at 90 degree increments.  This will provide equal resistance around the outside of the lens cell, which in turn will roughly center the lens.

- Place the dial indicator at the top edge of the lens and measure the wedge of the lens cell.  Adjust adding shim stock as necessary to level the lens to a .5 thou tolerance.

- Check the runout on the centration.  Adjust wedge and centration as necessary.  Record the final errors.

NOTE: Were not able to obtain the required runouts on the centration or wedge due to irregularities at the point of measurements, difficulty in room to add shims, and drag noticed with the dial indicator.

Error in final centration of the test lens 1: TIR 4.85 thou

Error in final wedge of the test lens 1: TIR 3.4 thou at diameter of 2 inches

- Mix the appropriate amount of Sylgard, evacuate the majority of bubbles.  Do not pump the sample for longer than 45 minutes.

- Hook up the automatic dispenser to the compressed nitrogen bottle to provide the appropriate pressure.  Using the dispenser, apply the Sylgard 184 layer carefully around the full edge of the lens with a continuous bead, slowly adjusting the pressure on the dispenser if the sylgard is dispensing too slowly or too swiftly.

- Allow the Sylgard to set.  It may take up to 4 days for the Sylgard to set sufficiently for moving the lens off of the table.  Once the RTV layer has set for at least 7 days, seal it with an oil resistant seal such as RTV 730 to seal all bond lines that may come in contact with the oil.

NOTE:  Instructions for the primer state that you should not let the primer set more than 6 hours prior to gluing, so primering of the cell should not occur before you are ready to align the lens in the cell and proceed immediately to the glue stage.

NOTE:  Encounters upon gluing.  Glue leaked out under lens.  Gap under lens less than 8 thou.  Sylgard did not cure where it came in contact with the shrink wrap.  After 5 days it was still very glue-like.

12. Mount test lens 2 using the same method as listed above.

Error in final centration of the test lens 2 cell: max to min 1.0 thou at 2.5 inches

Error in final wedge of the test lens 2 cell:  max to min 0.9 thou at 2.5 inches

Error in final centration of the test lens 2: 1.7 thou

Error in final wedge of the test lens 2: 0.5 thou at 2 inches

MOUNTING OF THE TEST CALCIUM FLOURIDE LENS:  A different process is used for this lens because there is not a ledge on the cell to rest the lens against.

- Measure the height of the outer adapter ring with the lens cell, at the top edge of the interior edge of the lens cell.  Next measure the height of the lens on the ball bearings, with the ring systems.  Compare the two measurements and add shims as necessary to make the two measurements equal.

- Place the standoff test ring, with internal ring, o-ring seal, and ball bearings on the rotation table, and roughly center it using a micrometer.  Lock it in place.

- Place the dial indicator at the outer edge of the raised alignment jig, and level the assembly to within +/- .5 thou. shimming as necessary.  See Figure XXXX.

- Next place the dial indicator at the outer edge of the alignment jig and center it to within +/- .5 thou.  Once this is accomplished, the test set-up will be clamped in place.

- Place the CaF2 lens cell on top of the jig and place the dial indicator on the top surface.  Measure the total indicator run out.  Wedge as necessary to level the cell.  Measure the height of the top lens mounting edge.

- Place the lens onto the set up, and roughly center it using a micrometer.  Place the dial indicator on to the top edge of the lens.  Measure the total run out.  Level the lens until it is within +/- .5 thou.  wedging under the o-ring as necessary.    Measure the height of the top edge of the lens, once the lens is roughly centered and flat.  This height will be verified at several of the positions around the outside edge of the lens, to check the wedge of the lens.

- Compare this height to the cell height.  If the cell height is too low, shim it higher, until the heights are the same.  If the lens height is too low, remove the lens and raise the inner ring until it is at the appropriate height for the lens.  Apply primer to the cell wall at this point.  Return the lens to the assembly and repeat steps 6 and 7 until the height of the lens and cell are correct for gluing, and have been properly leveled.

- Again, there is no mechanism to appropriately center the test lens, shim until it is roughly centered. Record the final wedge and decentration of the test CaF2 lens and cell:

Error in final centration of the test CaF2 cell:

Error in final wedge of the test CaF2 cell:

Error in final centration of the test CaF2 lens:

Error in final wedge of the test CaF2 lens:

- Apply the epoxy using the dispenser as outlined in step 10 of the previous section, until the appropriate thickness of the glue layer is obtained.  Verify that no wicking of the glue occurs past the o-ring seal.   Allow the glue to cure in place.  Once the glue has cured well (several days potentially), remove the test adapter plates and the o-ring seal.  Allow epoxy to cure for a minimum of 7 days.

#### ASSEMBLY OF THE TEST LENS TRIPLET

Once the individual lens have cured sufficiently, the test triplet assembly will be assembled as follows:

- Place the CaF2 lens assembly on the rotational table with the lens on the top side, and roughly center it using a micrometer.  Lock the cell in place.

- Place a dial indicator on the outer top edge of the cell and measure the runout.  Using shim stock, level the cell until it is within the tolerance.

- Place the dial indicator on the outer top edge of the cell, and center the cell until it is within tolerance.  Lock it in place.

- Verify the levelness of the cell again.  Adjust as necessary.

- Place the o-rings in place on the lens 1 cell.

- Place the lens 1 cell on top of the centered cell with the lens facing down, and roughly center it.

- Place the dial indicator on the top edge of the cell and measure the run out.  Do not shim the cell, as it could compromise the function of the o-ring seal.  Record the wedge.

Error in final centration of the test lens 1 cell in assembly:

Error in final wedge of the test lens 1 cell in assembly:

- Place the dial indicator on the outer edge of the cell and measure the total run out.  Adjust the centration of the lens cell until the total runout is within tolerance.  Lock into place.

- Lock the two lenses in place together.

- Flip the assembly over so that lens 1 is on the bottom, and lens 2 is on the top.  Roughly recenter the lens assembly on the table using the micrometer.  Lock in place.

- Place the dial indicator on the top of the outer edge of the assembly and measure the wedge.  Add shims under the whole assembly to eliminate this wedge.  Record final errors in decentration and wedge.

Error in final centration of the test lens 1 cell in assembly:

Error in final wedge of the test lens 1 cell in assembly:

- Place the next o-ring seal on the lens 2 cell.

- Place the test lens cell 2 onto the assembly with the lens facing down.

- Roughly center the lens 2 cell using a micrometer.  Place the dial indicator on the outer top surface of the cell and measure the total run-out on the cell. Do not try to level it at this point. As it may interfere with the operation of the o-rings.  Record the runout.

Error in final wedge of triplet assembly:

- Place the dial indicator on the outside edge of the top cell and measure the total runout.  Adjust the lens 2 cell until the appropriate tolerance is obtained.  Lock the cell into position with the other two lenses.  Record its decententration.

Error in final centration of the test lens 2 cell in assembly:

#### OIL CHECK OF ALIGNED SYSTEM

- Verify that the triplet is sealed.

- Install bleed off bladder. (Brake bleed?)

- Inject oil into the system, using the dispenser gun.

- Bleed off any air.

- verify system is filled with oil.

- Monitor the system for leaks.

You have now completed the alignment of the test optics and are ready to move on to the real thing.

MOUNTING OF LENSES IN CELLS

- Place a reference flat on the table.  This reference flat should have a diameter slightly less than the inner diameter of the lenses.  The exact diameter will vary with what is available from the mirror lab.  Roughly center the reference flat using a micrometer.

- Place the dial indicator on the very outer top edge of the reference flat, and measure the total runout.  Shim the mirror as necessary to ensure a level surface.

- Mount a laser above the table, using a jig that will allow for the mounting of the laser and a turning flat to allow for observing the reflection of the laser off the lower mirror by an alignment telescope mounted sideways.  See Figure 1.

- This will complete the set up of the alignment jigs.

#### MOUNTING OF THE SPARE LENS 3

Verify the full size lens alignment with the spare lens 3 assembly as follows:

- Place gauge blocks on the rotational table so they are equidistant from the center of the table.

- Place the lens cell for the spare lens 3 on gauge blocks, so that the lens may be placed in on the ledge. Roughly center the cell using a caliper.

- Place the dial indicator on the outer top surface of the cell and measure the total run-out of the top surface of the cell (Measurements every 20 degrees initially provides a good characterization).   Note the high and low spots and shim as necessary until the cell is level to well within tolerance. Adjust an outer edge holding jig so that you can use set screws to move the cell radially without interfering with the reference edge of the cell.  Adjust until it is barely holding the cell, but not so that it will deform the surface.  Check the levelness of the cell after holding in place.

- Move the dial indicator so that it resting against the reference surface on the outside edge of the cell.  Check the centration.  Using the set screws on the holding jig, adjust the lens until it is centered on the rotation of the table.  Lock the cell into place, so no accidental movement of the cell is possible.

- Verify the centration and levelness of the lens cell.  Adjust as necessary.

- When you are ready to begin aligning the lens 3 into the cell, paint the inside of the lens cell surface with the 92-023 primer.  Safety precautions should be taken. Once primer is applied, the lens must be aligned.  The manufacturer recommends that the adhesive be applied over the primer within 6 hours of primer application.  Tests performed by the Steward Observatory AO group indicate that overnight drying of the primer does not affect the performance of the adhesive strength.

- Place the appropriate precut lens pad into the cell and tape into place.

- Place the spare lens 3 into the locked down cell using the suction holder.  Place the lens so that the concave surface is facing up.  Roughly center the lens into the holder using shims.  Once roughly centered, remove the shims.

- Place the dial indicator on the top of the outer edge of the lens, and measure the levelness of the lens.  Shim under the lens until the lens is level.  Move the set screws in until they are barely resting against the lens edge but are applying no stress to the lens.

- Move the dial indicator to the outside edge of the lens, and measure the centration.  Using the set screws, slowly adjust the set screws to center the lens, but do not stress the lens or the cell.  Once the acceptable centration is achieved, adjust the set screw until they are stably holding the lens in place.

- Set another dial indicator on the top ledge of the lens and check the levelness of the lens.  Adjust the leveling shims as necessary.

- Direct the laser through the lens and monitor the positioning of the reflection using an alignment scope.  Monitor movement of the laser position on mirror as the table is rotated 360 degrees.  Adjust the centration and the wedge until there is no observable movement of the laser positioning.  Check the alignment with the dial indicators.

- Tighten the setscrews and verify the lens positioning again.  Repeat steps as necessary.  Remove temporary shims.

- Once the lens is securely locked in place, carefully apply putty to the edge of the cell shelf so that it does not come in contact with the OCA of the lens, but provides for a barrier to prevent leakage of the sylgard.

- Mix a SMALL amount of sylgard and apply in a continuous bead around the entire cell width.  Apply enough sylgard so that it can wick onto the shelf upon which the lens sits, but so that there is not enough to leak out the other side.  Allow to set sufficiently to aid in sealing the possible leak points.

- After the small amount of sylgard applied has set partially, mix remaining appropriate amount of Sylgard per directions, and evacuate bubbles.  Load dispenser with sylgard, and using the automatic dispenser apply the appropriate amount of Sylgard 184 in a continuous bead around the entire cell.  Continue evenly around cell to the desired sylgard width, and let set completely.

- Remove putty and clean surfaces affected.  Loosen set screws slightly so that no stress will be introduced to the lens.

- Mark the outside edge of the cell with an arrow to indicate the initial light path direction, and any remaining high or low points for the lens/cell combination, both in parallelism, and in decentration of the lens.

- Modify this procedure as necessary based on the obtained information of this test mounting.

- Continue mounting of the balance of lenses as per the following section.

LENS 2-6 MOUNTING PROCEDURE:

NOTE:  Element tilts for lenses 2 and 3 are critical with the tolerancing.  Make sure these alignments are performed extremely carefully and to the maximum possible accuracy.

- Place gauge blocks on the rotational table so they are equi-distant from the center of the table, and so that the lens 6 cell may be placed on the gauge blocks without interfering with the cut out.

- Place the lens cell lens 6 on gauge blocks, so that the lens may be placed in on the ledge. Roughly center the cell using a caliper.

- Place the dial indicator on the outer top surface of the cell and measure the total run-out on the cell.   Measure the high and low spots and shim as necessary until the cell is level to well within tolerance.

- Adjust an outer edge holding jig so that you can use set screws to move the cell radially without interfering with the reference edge of the cell.  Adjust until it is barely holding the cell.  Check the levelness of the cell after holding in place.

- Move the dial indicator so that it resting against the reference surface on the outside edge of the cell.  Check the centration.  Using the set screws on the holding jig, adjust the lens until it is centered on the rotation of the table.  Lock the cell into place.

- Verify the centration and levelness of the lens cell.  Adjust as necessary.

- When you are ready to begin aligning lens 6 into the cell, paint the inside of the lens cell surface with the 92-023 primer.

- Place the appropriate precut lens pad into the cell and tape into place, using two sided tape.

- Place lens 6 into the locked down cell using the suction holder.  Place the lens so that the concave surface is facing up.  Roughly center the lens into the holder using shims.  Once roughly centered, remove the shims.

- Place the dial indicator on the top of the outer edge of the lens, and measure the levelness of the lens.  Shim under the lens until the lens is level.  Move the set screws in until they are barely resting against the lens edge but are applying no stress to the lens.

- Move the dial indicator to the outside edge of the lens, and measure the centration.  Using the set screws, slowly adjust the set screws to center the lens, but do not stress the lens or the cell, until the acceptable centration is achieved.  Adjust the set screw until they are stably holding the lens in place.

- Set another dial indicator on the top ledge of the lens and check the levelness of the lens.  Adjust the leveling shims as necessary.

- Direct the laser through the lens and monitor the positioning of the reflection using an alignment scope.  Monitor movement of the laser position on mirror as the table is rotated 360 degrees.  Adjust the centration and the wedge until there is no observable movement of the laser positioning.

- Tighten the setscrews and verify the lens positioning again.  Repeat steps as necessary.  Once lens is locked into place, carefully remove temporary shims, and re-verify locked down position.  Carefully apply putty to the edge of the cell shelf so that it does not come in contact with the OCA of the lens, but provides for a barrier to prevent leakage of the sylgard.

- Mix a small amount of sylgard and apply in a continuous bead around the entire cell width.  Apply enough sylgard so that it can wick onto the shelf upon which the lens sits, but so that there is not enough to leak out the other side.  Allow to set sufficiently to aid in sealing the possible leak points.

- Mix Sylgard per directions, and evacuate bubbles.  Load dispenser with sylgard, and using the automatic dispenser apply the appropriate amount of Sylgard 184 in a continuous bead around the entire cell.  Continue evenly around cell to the desired sylgard width, and let set completely. (Minimum of 7 days)

- Remove putty and clean surfaces affected.  Loosen set screws slightly so that no stress will be introduced to the lens.

- Mark the outside edge of the cell with an arrow to indicate the initial light path direction.  In addition mark the low and high points of the decentration and wedge so that the best configuration can be obtained during the stack process.

- Mount lens 4 using this same procedure.  Lens 4 should be placed into the cell with the concave surface facing up.

- Mount lens 2 using this same procedure.  Lens 2 should be placed into the cell with the plano side up.

- Mount lens 3 using this same procedure.  Lens 3 should be placed into the cell with the concave surface up.

Error in final centration of lens 2 cell: max to min

Error in final wedge of lens 2 cell:  max to min

Error in final centration of lens 2:

Error in final wedge of lens 2:

Error in final centration of lens 3 cell: max to min

Error in final wedge of lens 3 cell:  max to min

Error in final centration of lens 3:

Error in final wedge of lens 3:

Error in final centration of lens 4 cell: max to min

Error in final wedge of lens 4 cell:  max to min

Error in final centration of lens 4:

Error in final wedge of lens 4:

Error in final centration of lens 6 cell: max to min

Error in final wedge of lens 6 cell:  max to min

Error in final centration of lens 6:

Error in final wedge of lens 6:

### MOUNTING OF LENS 5

The mounting of lens 5 will have to follow the same procedure as the mounting of the test Calcium Flouride lens since there will be no shelf for the lens to rest upon.

Note:  This section to be finalized after mounting of the test lens 5.

- Measure the height of the outer adapter ring with the lens cell, at the top edge of the interior edge of the lens cell.  Next measure the height of the lens on the ball bearings, with the ring systems.  Compare the two measurements and add shims as necessary to make the two measurements equal.

- Place the standoff test ring, with internal ring, o-ring seal, and ball bearings on the rotation table, and roughly center it using a micrometer.  Lock it in place.

- Place the dial indicator at the outer edge of the raised alignment jig, and level the assembly to within +/- .5 thou. shimming as necessary.  See Figure 2.

- Next place the dial indicator at the outer edge of the alignment jig and center it to within +/- .5 thou.  Once this is accomplished, the test set-up will be clamped in place.

- Place the CaF2 lens cell on top of the jig and place the dial indicator on the top surface.  Measure the total indicator run out.  Wedge as necessary to level the cell.  Measure the height of the top lens mounting edge.

- Place the lens onto the set up, and roughly center it using a micrometer.  Place the dial indicator on to the top edge of the lens.  Measure the total run out.  Level the lens until it is within +/- .5 thou.  wedging under the o-ring as necessary.    Measure the height of the top edge of the lens, once the lens is roughly centered and flat.  This height will be verified at several of the positions around the outside edge of the lens, to check the wedge of the lens.

- Compare this height to the cell height.  If the cell height is too low, shim it higher, until the heights are the same.  If the lens height is too low, remove the lens and raise the inner ring until it is at the appropriate height for the lens.  Apply primer to the cell wall at this point.  Return the lens to the assembly and repeat steps 6 and 7 until the height of the lens and cell are correct for gluing, and have been properly leveled.

- Again, there is no mechanism to appropriately center the test lens, shim until it is roughly centered. Record the final wedge and decentration of the test CaF2 lens and cell:

Error in final centration of Lens 5 cell:

Error in final wedge of Lens 5 cell:

Error in final centration of Lens 5:

Error in final wedge of Lens 5:

## Once sylgard has cured completely seal all surfaces that will come in contact with the sylgard 184 with a thin layer of RTV 730 sealant.

## The mounting of the lenses into their cells is now complete.  The next step is to start the stack process

## B.  START OF THE STACK PROCESS

- Schedule a block of time with the CMM machine at the Instrument lab.  A minimum of 2 months is desired.  Transfer the rotation table to this location to provide for double checks.

- Adjust the thermal control of the laboratory to 60 degrees, if possible.

- Place Lens 6 on the CMM table on gauge blocks, so that lens 5 can be stacked on top of it.  Fix lens 6 into its position on the CMM table.  This position will provide the main reference for stacking of the triplet. Do not move its position or re-home the CMM machine until the alignment of the triplet is complete.

- Quantify the positioning of the mount of lens 6, and the edge of the optic mounted inside.  Adjust as necessary to un-wedge any left over residual.  Fully characterize the positioning of both the lens and the cell., verifying that the plane of the edge of lens 6 is parallel to the plane of the CMM table, and that the optical axis defined by lens 6 is perpendicular to the CMM table.

- Once this lens has been fully characterized in its fixed position, place lens 5 on top of it with the appropriate spacer, and the smaller radius facing towards lens 6.

- Center lens 5 horizontally so that it is on optical axis with lens 6.

- Measure the position of the top surface of lens 5 and compare it to the calculated required position.  Check the measurements for wedge and spacing.

- Modify the spacer by approximately ¾ of the required correction to obtain the proper spacing.

- Remount lens 5 with the modified spacer.  Remeasure and repeat steps 7-8 until lens 5 is at the proper location, to within the acceptable tolerance.

- Direct a laser on axis through the elements and check for aberrations.  Adjust as necessary until the desired alignment is obtained.  Bolt elements together, maintaining alignment.

ACTUAL SPACER THICKNESS USED BETWEEN LENSES 5 AND 6:

Refer to actual as built drawing #??????

- Add lens 4 and the 4-5 spacer onto lens 5 with the curved surface facing lens 5.

- Repeat steps 7-8, comparing lens 4 position with lens 5 position.

- Once proper alignment has been obtained, bolt lens 4 together with lens 5-6.

- Check alignment with a laser through the optical axis and check for aberrations that may indicate decentration or wedging of the elements.

- Once satisfied with the alignment, make sure that the elements are well bolted so that no movement can occur with respect to each other.

- Turn triplet back onto the rotation table and set up with alignment telescope to monitor laser and see if decentrations are visible.

ACTUAL SPACER THICKNESS USED BETWEEN LENSES 4 AND 5:

Refer to as-built drawing # ?????

NOTE:  If the cut away of lens 6 interferes with the stacking on the table, we will reverse the order of the previous alignment.

NOTE:  Interferometric testing is not possible at this point without the design and fabrication of a holographic null lens.

LENSES 2 AND 3

NOTE:  Element tilts for lenses 2 and 3 are critical with the tolerancing.  Make sure these alignments are performed extremely carefully and to the maximum possible accuracy.

After Lenses 2 and 3 are mounted individually into their cells using the procedure listed in the previous portion of this document, they will be aligned to each other as follows:

- Place Lens 2 on gauge blocks on the CMM table so that the curved surface is facing down.  Measure the position of the lens and cell.  After shimming to define that the optical axis is perpendicular to the table, and the plane of the lens is parallel to the table, lock the lens into position so that it can become the reference position.  Again, do not move its position or re-home the CMM machine until the alignment of lens 2 to lens 3 is complete.

- Place lens 3 and the spacer between lens 2 and 3 on top of lens 2 with the concave surface facing down.  Adjust lens 3 horizontally until the optical axes of the two elements are in line with each other.  Measure the height and levelness of lens 3 and compare it to the required positioning.

- Modify the spacer by approximately ¾ of the required correction to obtain the proper spacing.

- Repeat steps 3 and 4 until the proper positioning of lens 3 is obtained.

- Direct a laser on axis through the elements and check for aberrations that would indicate some residual decentration or wedge.

- Bolt the two lenses together and verify alignment.

ACTUAL SPACER THICKNESS USED BETWEEN LENSES 2 AND 3:

See as-built drawing number ??????

REGISTRATION OF DOUBLET AND TRIPLET.

- Return lenses 2/3  to stack position on the CMM table with lens 2 down.

- Lens 3 will be held in position at this point, so that lens 2 and/or the triplet will be movable if necessary.  The position of lens 3 will be quantified at this point to be used as the reference point.

- Place the spacer for lens 3 to 4 and the triplet with lens 4 down on top of the doublet assembly and guide the triplet into position with push/pull screws, until it is concentric with the positioning of the lens 2/3 assembly, and the optical axes of the two subsections are coincident.

- At this point, the top of lens 4 will not be accessible, and spacing measurements will have to occur between lens 3 and lens 6.  Measure the spacing between lens 3 and 6.  Compare the positioning to the required spacing of these two elements.

- Adjustment the spacer thickness for the lens 3 to 4 spacer by ½ the required correction, and repeat the process to ensure that the appropriate spacing increment has been achieved.  Continue measurements and adjustment of the spacer until the desired spacing is obtained.

- Check the positioning of the lens assembly for perpendicularity of the lens optical axis with the CMM table and the concentricity of the elements.

- Lock the two subassemblies together and transfer it to the rotation table.  Center the system on the table.

- Observe the laser through the system with the alignment scope.  Verify the alignment and adjust as necessary.

- Perform a star test and interferometrically test lenses 2-6 to verify alignment.  Final adjustments will be made as necessary.  This test should be accomplishable with a precision flat to auto collimate the beam.  To properly do either of these tests, the oil must be added to the system.

LENS 1 ALIGNMENT

For the alignment of the cross roller slides to the cell assembly please refer to document number SP13879.  Below begins the alignment procedure for the installation of lens 1 to its cell, and the alignment of the completed assembly to the rest of the lens can.

- With the lens one cell on its side, as it will be mounted in the space frame, index mark 12 positions equally spaced around the outside of the inner cell for lens 1.  Using the CMM, measure these 12 positions.

- From these 12 positions, define the center position of the cell.

- At eight translation positions along the cross roller slides measure the positioning of the 12 points of the cell with the CMM machine.  Using the CMM define the movement centerline and error of this line.  This will become our optical axis of this assembly.

- Measure the front surface of the lens cell to define the plane.

- Measure the perpendicularity of the lens 1 faceplate plane and the optical axis plane.  Note any deviations.

- Lens 1 will have to be mounted perpendicular to the calculated optical axis, and on center with it.  Turn the assembly so that it is vertical.  Shim the assembly until the top surface of the inner lens 1 cell is perpendicular to the optical axis.  Using the CMM, shims as necessary and the side set screws, align lens 1 until it is in position with the optical axis of the lens 1 assembly to within the tolerances specified.

- It is anticipated that the way we have specified the mechanical mounting of the cross roller slides, that the front surface of lens 1 should be perpendicular to the optical axis, and very little shimming should be necessary between lens 1 and lens 1 cell.

- Once the position of lens 1 is correct in the cell and is firmly being held in place with the setscrews, turn the lens over and apply a very fine layer of putty against the shelf, but outside of the clear aperture of lens 1.  This will prevent wicking of the sylgard onto the surface of lens 1 if shims are too great.

- Turn the lens back over and reverify the positioning of the lens.  Adjust if necessary.

- Mix a small amount of sylgard and apply it in a thin bead around the cell.  This will be used to wick through the shelf and provide an additional barrier to the full sylgard amount.

- Once the first amount of sylgard has set so it will no longer flow, mix the remaining amount of sylgard and apply it to fix the lens within the cell.    While the sylgard is setting, it is important not to move the lens.  It may take up to 4 days for the sylgard to harden to the point that the lens will be movable from its location.

- Fix the lens 2-6 assembly in place on the rotation table with lens 6 down.  Next, attach the spool assembly that connects the lens 1 assembly to lenses 2-6.  Center this with the lens 2-6 assembly to within the mechanical tolerances of the spool assembly.

- Attach the lens 1 assembly and center the inner cell with respect to the lens 2-6 assembly using the push-pull tip-tilt screws on the lens 1 assembly, while monitoring a dial indicator that is recessed down to the edge of the lens 1 inner cell. Observe a laser with the alignment scope, and make adjustments until there is no wobble of the laser viewed through the lens system, and the dial indicator indicates they are concentric and level.  Check the position of lens one on the cross roller slides to ensure that the position is centered on the available movement of the lens.

- Transfer the entire set up back onto the CMM machine, and measure the distance between lens 1 and lens 6.  Compare this to the required separation, and adjust the spool size until the desired separation is achieved, and the lens is still near it’s center of movement.

- Fix the full collimator together.

- Using a star target, monitor through an alignment telescope to verify the centration of the full collimator.  Adjust with the push-pull tip-tilt screws on the lens 1 assembly until the system is aligned.

- Move the entire assembly to an interferometer, and check the image quality.  Adjust lens 1 if necessary to improve image quality.  If no improvement is obtainable, and the image quality is poor, earlier steps of this procedure will have to be repeated, until collimator is fully aligned to optimum performance.

- Continue on the alignment of the system using procedures:

SP14098- Prism alignment procedure

SP14099- System alignment procedure

SP????- Space frame construction

Sp???- -  Grating alignment procedure